## Slide 1
Accelerators for ion beam therapy
MPA Retreat 2024
30.09.2024
Christian Schömers

## Slide 2: Outline
Introduction
  Overview of therapy accelerators in the world
  Rationales for particle therapy (PT): biological / physical / technical aspects
  Clinical examples
The accelerator
  Cyclotron and examples
  Synchrotron, example facility in Heidelberg: HIT
  Comparison to other accelerators for PT in the world
Recent developments in PT (very limited and personal choice)

## Slide 3: Outline
Introduction
  Overview of therapy accelerators in the world
  Rationales for particle therapy (PT): biological / physical / technical aspects
  Clinical examples
The accelerator
  Cyclotron and examples
  Synchrotron, example facility in Heidelberg: HIT
  Comparison to other accelerators for PT in the world
Recent developments in PT (very limited and personal choice)

## Slide 4: What are medical facilities? In the accelerator context
Particle Therapy Centers doing patient treatment with protons or heavier ions
109 Protons
15 Carbon (+ p)
ptcog.ch/ (Sept. 2024) Facilities in operation only.
48
30
28
2

## Slide 5: Rationales for proton and ion beam therapy
Krämer & Durante, Eur. Phys. J. D 60 (2010)
Depth-dose profile: Bethe-Bloch
Range defined by starting energy (Bragg peak)
Higher biological effectiveness
Less lateral scattering
Possibility of lateral guidance (scanning)

## Slide 6: Rationales for proton and ion beam therapy
Amaldi & Kraft, Rep. Prog. Phys. 68 (2005)
Depth-dose profile: Bethe-Bloch
Range defined by starting energy (Bragg peak)
Higher biological effectiveness
Less lateral scattering
Possibility of lateral guidance (scanning)

## Slide 7: Rationales for proton and ion beam therapy
Depth-dose profile: Bethe-Bloch
Range defined by starting energy (Bragg peak)
Higher biological effectiveness
Less lateral scattering
Possibility of lateral guidance (scanning)
Protons
Carbon ions
ca. 30 cm
ca. 30 cm

## Slide 8: Rationales for proton and ion beam therapy
Depth-dose profile: Bethe-Bloch
Range defined by starting energy (Bragg peak)
Higher biological effectiveness
Less lateral scattering
Possibility of lateral guidance (scanning)
Dose-delivery system
Haberer et al., NIM A 330 (1993)

## Slide 9: How we do therapy State of the art: raster-scanning
Pencil beam is
guided by
  scanning magnets
monitored by
  multi-wire chambers
  ionization chambers
Spread out Bragg Peak
Accelerator delivers a spectrum of ion energies
courtesy of Siemens Health Care

## Slide 10: Clinical examples Photons vs. ions I
Conventional (photons, 9 fields)
Charged particles (2 fields)
Kraemer M and Scholz M, Treatment planning for heavy ion therapy, Phys. Med. Biol. 45 3319–30

## Slide 11: Further examples of tumours treated with ions Skull
Chordoma and chondro-sarcoma at the base of the skull, Adenoid cystic carcinoma in the head and neck region (salivary glands)
Characteristics:
  Deep located
  Radio-resistant
  Close vicinity to organs at risk (optical nerves, brain stem)

## Slide 12: Clinical examples Photons vs. ions II
Conventional (photons)
Charged particles
Courtesy of K. Herfarth

## Slide 13: What makes a medical facility special?
“Beam target” is a patient!
Mission: provide particle beams in therapy quality 335 days / year
Avoid any interruptions to guarantee therapy success
The dose delivery must be safe and precise
Fast: reduction of irradiation time only without affecting safety and precision
Large variety of beam parameters (ion types, energy steps, spot sizes, intensity range, gantry angles)

## Slide 14
Introduction
  Overview of therapy accelerators in the world
  Rationales for particle therapy (PT): biological / physical / technical aspects
  Clinical examples
The accelerator
  Cyclotron and examples
  Synchrotron, example facility in Heidelberg: HIT
  Comparison to other accelerators for PT in the world
Recent developments in PT (very limited and personal choice)

## Slide 15: Beam requirements
Energy defined by required penetration depth:
For 30 cm:
  220 MeV for p
  430 MeV/u for 12C6+
Intensity:
  few ~ 109 protons/s
  few ~ 107 12C6+/s

## Slide 16: Cyclotrons
D-shaped RF electrodes (“Dees”) placed in a disk-like vacuum chamber and embedded in a large (near-homogeneous) static magnetic field.
Radius of particle trajectory increases at each passage through the gap.
But the cyclotron-frequency is constant (non relativistic case)!
Remark: Besides Proton-beam therapy, cyclotrons are also used as Radioisotope factories for nuclear medicine. → The cyclotron is the most used ion-accelerator for medical purposes.
E. Lawrence's original concept of the cyclotron (1934 patent)

## Slide 17: Cyclotrons for ion-beam therapy
C230 by IBA
Designed for proton therapy
Installed at 16 facilities.
Mass: 220 t
B = 2.2 T
A. Peters, CERN Academic Training Lecture (2013)

## Slide 18: Cyclotrons for ion-beam therapy
www.psi.ch
COMET superconducting cyclotron
Developed by ACCEL / PSI (later Varian Medical) for proton therapy
Mass: 80 t
Bmax = 3.8 T

## Slide 19: Cyclotrons for ion-beam therapy: Energy selection
A. Peters, CERN Academic Training Lecture (2013)
beam diag.
degrader
collimator
Multi-wedge degrader (PSI): 238 … 70 MeV
250
Accelerator: 250 MeV
Energy-selecting slits
A cyclotron is a fixed energy machine
→ Degrader + momentum-selection
At low final E >99% of all p are lost! → Strong activation of beamline components.

## Slide 20: Cyclotrons for ion-beam therapy: 12C6+ ?
“magnetic rigidity”: Relates particle momentum and charge to magnetic field and bending radius
Protons (220 MeV): Bρ = 2.3 Tm
12C6+ (430 MeV/u): Bρ = 6.6 Tm
To go from protons to carbon ions, one needs to increase either the magnetic field or the size of the machine by a factor ~3.
Although they are very successful in proton therapy, there is no cyclotron for carbon-ion therapy yet.
But: Project ARCHADE to install a carbon-treatment facility in Caen (France).
Based on a superconducting cyclotron (“C400”) developed by IBA.
The C400: the ARCHADE multi-ions cyclotron, Jacques BALOSSO, Workshop on Ion Beam Therapy,
Archamps, 2018

## Slide 21: Synchrotrons
Linac
Deflection magnets (dipoles)
Accelerating cavity (fRF, URF)
Focusing magnets (quadrupoles)
Extraction beam line
Beam
injection
Beam
extraction
Medium energy ions
High energy ions
circulating ions
Advantage: Individual elements (magnets, vacuum
chambers) can be
small …
… but the size of the machine is almost unlimited. Typical size of p/C therapy synchrotron: 20 m diameter

## Slide 22: Cyclotron beam vs. Synchrotron beam
Pros
Scalable, can deliver heavy ions.
Variable energy, (almost) no particle losses.
Easy to operate different ion species
Cons
Pulsed operation, breaks between cycles
Very high dose rates hard to achieve
Pros
Continuous train of short particle bunches.
High intensity
Cons
Fixed energy: needs degraders (particle loss, radiation safety issues).
Heavy ions (12C6+) would require very large magnet and chamber.

## Slide 23
Introduction
  Overview of therapy accelerators in the world
  Rationales for particle therapy (PT): biological / physical / technical aspects
  Clinical examples
The accelerator
  Cyclotron and examples
  Synchrotron, example facility in Heidelberg: HIT
  Comparison to other accelerators for PT in the world
Recent developments in PT (very limited and personal choice)

## Slide 25: HIT History – beam parameters – patients
Successor of GSI pilot project (440 patients treated)
1st hospital based particle therapy facility in Europe
≈ 700 patients/year since 2009
Total No. of patients treated > 9000
Treatment with:
  Carbon ions
  Protons
  Helium ions: Regular treatments since July 2024
  Oxygen ions: only for experiments
255 energy steps each (penetration depth: 2cm…30cm)
Different spot sizes
Different intensity levels
fair-center.eu

## Slide 26: Facility design Accelerator layout
Treatment rooms
Synchrotron,
(48 - 430 MeV/u)
Beam transport Beam shaping
Abort systems
Sources
Injector LINAC
Experimental room

## Slide 27: Ion Sources
ECR ion source (Electron Cyclotron Resonance)
Generation of ion plasma through microwave heating
Gas
RF
Permanent magnets

## Slide 28: Injector-LINAC (LINear ACcelerator)
Why we need a LINAC:
  Small beam energy of the ECR ion source implies extremely low B-field in the synchrotron
  Intensity of C6+ from ECR not sufficient → must extract C4+ and strip ions later

## Slide 29: What is part of the LINAC? RFQ (Radio-Frequency-Quadrupole)
0,25 m
Sinusoidal structure of 4 electrodes
8keV/u → 400keV/u
An RFQ can do 3 things in parallel:
Transverse focusing
Bunching
Acceleration

## Slide 30: What is part of the LINAC? IH Drift Tube Linac
Acceleration of the beam particles
Drift of the beam particles during field reversal
Drift tubes
Additional focusing elements
400keV/u → 7MeV/u

## Slide 31: Medium Energy Beam Transport (MEBT)
Function
Transport the beam to the synchrotron
Strip ions and charge separation of stripped ions
Optimisation for good injection conditions (longitudinal and transversal)
Foil stripper: 100 µg/cm² Carbon
H3+
4He2+
12C4+
16O6+
P
4He2+
12C6+
16O8+

## Slide 32: Synchrotron
Beam injection and accumulation
Acceleration to the final energy
Slow extraction allowing for beam spills of several seconds with interruptions
Setup:
6 dipoles 60°
12 Quadrupoles
4 Sextupoles
2 RF cavities
RF-knockout exciter
…

## Slide 33: High Energy Beam Transport (HEBT)
Beam transport to treatment rooms
Spill abort systems
Quadrupoles for final focusing
Dipoles and Steerer for position correction
Scanning magnets

## Slide 34: Horizontal Treatment Rooms
Beam application and monitoring system
Imaging system
Patient table
Laser positioning system
Patient treatment requires preparation and post-processing time
(Patient positioning, imaging, …)
 To optimally use the accelerator, we need more than one treatment room

## Slide 35: Heavy-ion Gantry
beam
First heavy ion gantry worldwide
Arbitrary entrance angle
Full 360° rotation
Overall weight > 600 tons
Diameter: 13m
Length: 25m
In operation (for treatments) since 2012.
Accuracy of beam position in isocenter:  0.5 mm

## Slide 36: Experimental Room
Dedicated to R & D
Equipment similar to treatment rooms
Even higher flexibility of beam parameters

## Slide 37: Typical therapy day @HIT
4h 6h 8h 10h 12h 14h 16h 18h 20h 22h 24h 2h 4h
Patient treatment
Med. Phys.QA
Acc QA
R & D
or
Beam adjustment
QA,
Plan Verific.
Accelerator QA (1h):
  Measurements along the accelerator
  Beam size and position
  Particle rates and spill structure
  Beam currents in LEBT-LINAC-Section
  Output energy of LINAC

## Slide 38: Typical therapy day @HIT
4h 6h 8h 10h 12h 14h 16h 18h 20h 22h 24h 2h 4h
Patient treatment
Med. Phys.QA
Acc QA
R & D
or
Beam adjustment
QA,
Plan Verific.
Medical Physics QA (2h):
  Room certification
  Position and width check
  Beam abort systems
  Patient table movement
  Treatment plan verification

## Slide 39: Typical therapy day @HIT
4h 6h 8h 10h 12h 14h 16h 18h 20h 22h 24h 2h 4h
Patient treatment
Med. Phys.QA
Acc QA
R & D
or
Beam adjustment
H1
H2
Gantry
QA,
Plan Verific.
Room activity (15th Nov. 2018) 71 patients

## Slide 40: Typical therapy day @HIT
4h 6h 8h 10h 12h 14h 16h 18h 20h 22h 24h 2h 4h
Patient treatment
Med. Phys.QA
Acc QA
R & D
or
Beam adjustment
QA,
Plan Verific.
R & D
  Bio-physics experiments
  Detector development
  Accelerator R & D
Accelerator beam adjustment

## Slide 41: Other particle therapy facilities
Introduction
  Overview of therapy accelerators in the world
  Rationales for particle therapy (PT): biological / physical / technical aspects
  Clinical examples
The accelerator
  Cyclotron and examples
  Synchrotron, example facility in Heidelberg: HIT
  Comparison to other accelerators for PT in the world
Recent developments in PT (very limited and personal choice)

## Slide 42: Synchrotron ion-beam therapy facilities around the world
Example Chiba, Japan:
HIMAC (Heavy-Ion Medical Accelerator in Chiba, NIRS, 1994)
Two 800 MeV/u synchrotrons, for ions up to 40Ar18+, mostly 12C6+.
Example: CNAO, Pavia, Italy, cortesy of M. Pullia
National Center for Oncological Hadrontherapy
Ion Source
Linac
Synchrotron
Room E
S.C.
Gantry
Room F
HIMAC
New Facility
Room
G
(under construction)
Room
B
Room
C
Room
A
cortesy of K. Noda

## Slide 43: Synchrotron ion-beam therapy facilities around the world
MIT (Marburg) Layout, cortesy of C. Krantz
MedAustron (Wiener Neustadt, Austria) Layout, courtesy of M. Kronberger

## Slide 44: Synchrotron proton-beam therapy facilities
1st hospital-based proton therapy centre (since 1990)
Proton synchrotron designed by Fermilab
> 20000 patients treated
Loma Linda, USA
Endo, M. Radiol Phys Technol (2018) 11: 1

## Slide 45
Introduction
  Overview of therapy accelerators in the world
  Rationales for particle therapy (PT): biological / physical / technical aspects
  Clinical examples
The accelerator
  Cyclotron and examples
  Synchrotron, example facility in Heidelberg: HIT
  Comparison to other accelerators for PT in the world
Recent developments in PT (very limited and personal choice)

## Slide 46: Reduce size and costs Cyclotrons - Single-room solutions
Proteus ONE single-room solution by IBA:
40 installations world-wide
iba-worldwide.com
S250 system by Mevion
Gantry-mounted superconducting cyclotron.
Installed at 15 sites.
mevion.com

## Slide 47: Reduce size and costs Synchrotrons - single-room solutions
Umezava et al., Hitachi Rev. 64 (2015) 508
Hitachi ProBeat V System
Compact p-synchrotron (250 MeV) with single-room solution
hitachi.com
ca. 5 m

## Slide 48: Reduce size and costs Synchrotrons – State of the art
Toshiba Heavy Ion (12C6+) Therapy System
Installed at multiple sites in Japan and South Korea.
global.toshiba
www.toshiba.com/heavy-ion/

## Slide 49: “Plan of Miniaturizing”, the Japanese way
10 m
10 m
1st Generation Quantum Scalpel
2nd generation Quantum Scalpel
23 m
10 m
4th generation
3rd generation
Ion Therapy in Japan
Koji Noda
National Institute of
Radiological Sciences QST Workshop on Ion Beam Therapy,
Archamps, 2018

## Slide 50: HIT – current projects
Control system upgrade to allow for multi energy operation
New RFQ
H. Höltermann, H. Hähnel, B. Koubek, H. Podlech, U. Ratzinger TUNING AND RF MEASUREMENTS OF THE LILAC RFQ Proc. of 15th Int. Conf. on Heavy Ion Acc. Technology HIAT2022, Darmstadt, Germany
IES #1
IES #2
IES #3
B
t
t
IES #1
IES #2
IES #3
B
Phases without beam availability at target (Injection, acceleration, reset, …)
C. Schoemers et al.
"Reacceleration of ion beams for particle therapy."
Proc. of IPAC2014, Dresden.

## Slide 51: “FLASH-effect” – Ultra High Dose Rate (UHDR) Irradiations
What is uHDR: dose rates > 40Gy/s (while standard dose rates (SDR) are << 1Gy/s)
Under investigation: are there biologic benefits when applying uHDR?
“FLASH effect”: normal tissue sparing without sacrificing tumour control?
Example helium: 3.5e10 particles in 90ms  3.9e10 part./s. > 50 times higher than therapy
T. Tessonnier, S. Mein, D.W.M. Walsh, N. Schuhmacher, H. Liew, U. Weber, R. Cee, M. Galonska, C. Schömers, S. Scheloske, S. Brons, J. Debus, T. Haberer, A. Abdollahi, A. Mairani, and I. Dokic. FLASH dose-rate helium ion beams in vitro. International Journal of Radiation Oncology, Biology, Physics, 2021.
C. Schömers, S. Brons, R. Cee, A. Peters, S. Scheloske, T. Haberer, “BEAM PROPERTIES BEYOND THE THERAPEUTIC RANGE AT HIT”, in Proc. IPAC’23, Venice, Italy, May 2023, paper THPM064, http://doi/jacow-ipac2023-thpm064/index.html

## Slide 52: iRAD – Ion Beam Radiography
Idea of ion beam radiography: use particle beams also for imaging
Good for treatment planning – same radiation for image and treatment
Possibility of in situ monitoring
T. Gehrke et al., “Theoretical and experimental comparison of proton and helium-beam radiography using silicon pixel detectors”, Phys. Med. Biol., vol. 63, no. 3, p. 035037, Feb. 2018. https://doi:10.1088/1361-6560/aaa60f
A. Burker et al., “Imaging with Ion Beams at MedAustron”, Nuclear Inst. and Methods in Physics Research, A 958 (2020) 162246, https://doi.org/10.1016/j.nima.2019.05.087
Object /
Patient
Detector

## Slide 53: iRAD – Ion Beam Radiography Carbon – Helium mixed beam
L. Volz, L. Kelleter, S. Brons, L. Burigo, C. Graeff, N. I. Niebuhr, R. Radogna, S. Scheloske, C. Schömers, S. Jolly, and J. Seco. Experimental exploration of a mixed helium/carbon beam for online treatment monitoring in carbon ion beam therapy. Physics in Medicine & Biology, 65(5):055002, feb 2020.
C. Knobloch, M. Metzner, F. Kehrein, C. Schömers, S. Scheloske, S. Brons, R. Hermann, A. Peters, O. Jäkel, M. Martisikova and T. Gehrke. Establishment and evaluation of a water-equivalent thickness calibration for helium-beam radiography with high-energy beams and thin silicon pixel detectors. Medical Physics, to be published 2022
Treatment and online monitoring at the same time
Accelerate 2 ions with same a/q to same energy
Carbon ions stop in the patient
Helium ions stop in detector
Simultaneous extraction!

## Slide 54: Multi ion treatment
What is multi ion treatment: treating the patient with more than one ion type.
Why? Maximize the therapeutic effects; make use of different LETs (=linear energy transfer, “severeness” of damage)
Two ions  twice the number of slices  double irradiation time
To be comparably fast as now: increase the speed of dose application e.g. higher intensity.
Taku Inaniwa et al, Adaptation of stochastic microdosimetric kinetic model to hypoxia for hypo-fractionated multi-ion therapy treatment planning, 2021 Phys. Med. Biol. 66 205007

## Slide 55: Particle ARC therapy
What is ARC-therapy: delivery technique with dosimetric advantages (dose conformity, treatment time)
Live/dynamic gantry rotation during irradiation instead of 2-3 fields “step and shoot”.
Already used in standard photon therapy  What about particle-therapy: spot-scanning hadron arc (SHArc)?
beam
Stewart Mein, Thomas Tessonnier, Benedikt Kopp, Christian Schömers, Semi Harrabi, Amir Abdollahi, Jürgen Debus, Thomas Haberer, Andrea Mairani, Biological Dose Optimization for Particle Arc Therapy Using Helium and Carbon Ions International Journal of Radiation Oncology*Biology*Physics, Volume 114, Issue 2, 2022, Pages 334-348, https://doi.org/10.1016/j.ijrobp.2022.04.025

## Slide 56: Take-home messages
Introduction
  Overview of therapy accelerators in the world
  Rationales for particle therapy (PT): biological / physical / technical aspects
  Clinical examples
The accelerator
  Cyclotron and examples
  Synchrotron, example facility in Heidelberg: HIT
  Comparison to other accelerators for PT in the world
Recent developments in PT (very limited and personal choice)

## Slide 57
…and all people for the discussions, that gave input and contributed to this talk, especially:
HIT: Thomas Tessonnier, Malte Ellerbrock, Rainer Cee, Andreas Peters, Thomas Haberer and the entire accelerator team Siemens Health Systems: Abel Robin GSI: Claude Krantz, Michael Galonska MIT: Alexander Schottelius CNAO: Marco Pullia MedAustron: Matthias Kronberger-Mayerhofer
HIMAC: Koji Noda
Thank you!

## Slide 58: Backup-Slides

## Slide 59: Spillpause – mechanism to spare healthy tissue
Spill pause is a fast beam-off mechanism
Pause and resume the beam extraction several times on flattop
RF-KO allows spill pause
Used for:
  Scanning without dose deposition
  Respiratory gating
Irradiation window
Gating
Target movement
Time
Courtesy of S. Nill & U. Oelfke, DKFZ, C. Bert et al
Tumour
Scan path

## Slide 60: HIT – Operating concept
Accelerator team: ~20 people
  Physicists, engineers, technicians
  2 accelerator operators on-site, 24/7
  On-call duty for each maintenance group
Availability of accelerator:
  > 330 days of operation per year
  1 shutdown Christmas – 1st week in January
  5 maintenance blocks 2 days each
  12 single maintenance days

## Slide 61: Requirements from dose distribution techniques
Durante and Paganetti, Rep. Prog. Phys. 79 (2016)
State of the art:
Pencil-beam scanning
”Paint” iso-energetic slice of the
target volume.
Requires beam size, energy, and intensity to be stable over several seconds!

## Slide 62: Topic 1: Intensity control system of a slowly extracted ion beam
Acceleration
Injection
Extraction
Reset
B
t
Notes: The first topic now is the intensity control system that has been fully implemented in the therapy operation yet
Let's have a look on the standard synchrotron cycle that somehow looks like this
When you try to shorten the overall cycle you can
  either try to reduce the time, while the patient receives dose
  or try to reduce the acceleration and preparation phases
  We have examples for both in this talk
This development deals only with the extraction phase

## Slide 63: Intensity control system
Use intensity signal
Add feedback loop
Position and intensity is measured, the scanning velocity and the intensity are adapted
Therapy- Control- System
Intensity feedback loop
RF-KO exciter

## Slide 64: Topic 1: Intensity control system Setup "Dynamic Intensity Control"
Intensity is monitored →
  Scanning velocity is adapted
  AND
  Beam intensity is adapted
Fast response due to RF-KO extraction (~ms)!
Therapy control system
Intensity feedback
"DIC"
RF-knockout
exciter
Notes: This is again visualized here:
  The beam is measured in front of the patient with ICs and MWPCs
  The therapy control system adapts the scanning magnets
What we did is we closed another feedback loop
We take the signal of the ionisation chambers, that monitor the beam intensity anyway
Instead of just adapting the scanning speed, we now directly also influence the Amplitude of the transverse KO-Exciter and thus delivered beam intensity
This is possible, as the knockout exciter can adapt the intensity on a ms timescale.
So we here coupled the industrial product, the accelerator and the medical product, the scanning system

## Slide 65: Extraction system: RF knock-out
More Details about the particular excitation, frequencies, hardware and recent improvements achieved in the spill structure in the talks of Eike Feldmeier and Cristopher Cortés this afternoon.
Ideal - flat
Typical RF-KO
Intensity controlled
RF-Knockout-Exciter parameters
Horizontal resonance tune: 1.67
RF-KO electrode length: 0.7 m
RF-KO electrode distance: 0.15 m
Maximum exciter power: 400 W
Intensity feedback system implemented (dose delivery system ↔ RF-KO exciter)

## Slide 66: Topic 1: Intensity control system Spill with intensity control
Intensity is monitored →
  Scanning velocity AND
  Beam intensity is adapted
Typical spill shape at HIT after April 2013
  With feedback
  Reference value
  Without feedback
Higher average intensity
Less treatment time
Notes: This is the spill shape after we closed the feedback loop, compared to the spill you have seen before
We can keep the intensity close to the desired value
The average is higher due to several reasons
  As there are less fluctuations, we can go closer to the upper limit, without having the risk of irradiating to fast
  Fast rise time, which is significant, as many spills do not last 5s but maybe just one or two, in case the slice is already completed
We still had, however, one fixed intensity for the whole cycle;
rasterpoints that require only a low dose, are irradiated with the same intensity as very "large", high-dose rasterpoints, which than take much longer to be completed

## Slide 67: Why a raster-point-individual intensity makes sense …if you have an intensity control
Dose in each voxel ≠ dose per raster-point
Example of dose distribution of one slice
Fixed intensity: irradiation time per raster point can vary by a factor of 2000!
ions
C. Schoemers et al, Nuclear Instruments and Methods in Physics Research Section A: Accelerators, Spectrometers, Detectors and Associated Equipment, Volume 795, 21 September 2015, Pages 92-99, ISSN 0168-9002, http://dx.doi.org/10.1016/j.nima.2015.05.054

## Slide 68: Topic 3: Multi-energy operation of a synchrotron
Acceleration
Injection
Extraction
Reset
B
t
Notes: The third and last topic of cycle time reduction is what we call multi-energy operation
It focuses mostly on the acceleration and reset phase

## Slide 69: Topic 3: Multi-energy operation Present situation
Energy #1
Energy #2
∆E < 4 MeV/u
Extr = 1.5 s
(5 s possible)
B
t
C. Schoemers et al.
"Reacceleration of ion beams for particle therapy."
Proc. of IPAC2014, Dresden.
HIMAC approach: Y. Iwata et al.
"MULTIPLE-ENERGY OPERATION WITH QUASI-DC EXTENSION OF
FLATTOPS AT HIMAC."
Proc. of IPAC10, Kyoto.
Notes: Today, each iso-energy-slice (IES) requires new synchrotron cycle.
Remaining particles at the end of extraction are dumped
and usually there are still particles left after completing one iso energy slice:
At HIT the average extraction time in therapy is 1.5 seconds, usually we have enough particles for 5s.
In the typical therapy case, the next cycle has just a slightly different energy of some MeV/u
We put a lot of effort of ramping down, ramping up again to end up with an energy that is almost the same.
Our Japanese colleagues realized the potential and started a very interesting work already some years ago and they published several papers on that
they have chosen the approach of having a pattern of the whole available energy range and stop this pattern, to have an extraction flattop whenever needed

## Slide 70: Topic 3: Multi-energy operation Concept
E #1
E #2
E #n
//
B
t
Accelerate (or decelerate) n times according to the patient-specific treatment plan
Iso energy slices irradiated with short interruptions
HIT test system with 2 extraction flattops available now
M. Galonska et al. "Multi-Energy Trial Operation of the HIT Medical Synchrotron: Accelerator Model and Data Supply."
Proc. of IPAC2017, Copenhagen.
E #n-1
E #n-2
Notes: Our approach is similar but different
Fill synchrotron to maximum
Accelerate or decelerate remaining particles to the next energy level
but we want to do this again in a patient specific way, depending on the demands of the actual treatment plan
this means at any time in the cycle, if possible to any required energy
Several iso energy slices can be irradiated with short interruptions only
as long, as there are still some particles left in the synchrotron
in average we will be able to have about 6 extraction levels per cycle
in some cases one cycle will even be enough to cover the whole tumour
We changed our data supply model in a way, to be able to have 2 extraction levels within one cycle, as this was possible with only moderate changes of the existing control system hardware

## Slide 71: Topic 3: Multi-energy Test-system results: Spill
Interspilltime < 300ms (TCS dominated)
Energy verified ü (PTW Peakfinder water column)
Next step: control system upgrade!
1st Extraction @
E1 = 352 MeV/u
2nd Extraction @
E2 = 358 MeV/u
Notes: You clearly see 2 spills on the ionization chamber of the HEBT-line
With a very short interruption in the order of some 100ms, which is btw. not dominated by the accelerator, but some internal communication of the therapy control system
We verified the energy in the treatment room with the standard device for range measurements at HIT
The next big step will be to think about a larger control system update and a new data supply model, that allows energy jumps according to the individual treatment plan

## Slide 72: Particle ARC therapy What the accelerator has to provide
Intensity range comparable to today
But: Stability of intensity!
Apart from SX: Multi-energy mode  fast energy switching mandatory, more degrees of freedom, robust optimization, better coverage, faster irradiation time!
Multi-energy mode requires beam off / on spill pause mechanism